## Sheet1
Test Step | ERC Test ID | Measured Parameter description | Unit of measure | LSL | USL | How to record in EDHR
Inspection | 15001 | FRU crate visual inspection & packaging correctness | pass/fail | NA | NA | Manullement
| 15002 | Installability checks | pass/fail | NA | NA
| 15003 | Labeling verification | pass/fail | NA | NA
Diagnostic | NA | SN, PN, Operation and? |  |  |  | ATE
Output conditionnal tunning | 14001 | Head/Body RF Port Isolation | dB | 30 | - | Manuellemnt ou récupérer les résulats sur un banc à côté
| 14002 | Body RF Sample Outputs A gain | dB | -51 | -49
| 14003 | Body RF Sample Outputs B gain | dB | -51 | -49
| 14004 | Head RF Sample Outputs A gain | dB | -41 | -39
| 14005 | Head RF Sample Outputs A gain | dB | -41 | -39
| 14011 | TR Input/RF Output Port Isolation | dB | 35 | -
| 14009 | Head TR resistance to head output | ohm | 0 | 0.1 | Manuellemnt, mesuré par un multimètre
| 14010 | Body TR resistance to Body output | ohm | 0 | 0.1
Power module gain tunning | NA | Ajuste the gain of master and slave modules by sw | pass/fail |  |  | ATE
Input conditionnal board tunning | 12001 | Body RF port; RF In -4dBm +/-0.25 dBm, into 50ohms, ajust the gain to get Body output to PEP | dBm | 72.04 | - | Manullement
| 12002 | Body RF port; RF In -4dBm +/-0.25 dBm, into 50ohms, ajust the gain to get Body output to PEP | dBm | 61.04 | 65.04
| 12003 | Body RF port; RF In 10dBm +/-0.25 dBm, into 50ohms, ajust the gain to get Body output to PEP | dBm | 72.04 | -
| 12004 | Body RF port; RF In 10dBm +/-0.25 dBm, into 50ohms, ajust the gain to get Body output to PEP | dBm | 61.04 | 65.04
| 12005 | Body RF port; RF In 3.5dBm +/-0.25 dBm, into 50ohms, ajust the gain to get Body output to PEP | dBm | 72.04 | -
| 12006 | HEAD RF port; RF In 3.5dBm +/-0.25 dBm, into 50ohms, ajust the gain to get heas output to PEP | dBm | 61.04 | 65.04
performance test - burn | 12007 | Input VSWR < 1.5:1 | - | - | 1.5 | ATE
| 13101 | Body Bandwidth (±275kHz) | dB | 0 | 0.4
| 13102 | Head Bandwidth (±275kHz) | dB | 0 | 0.4
| 13103 | Body output power nominal | dBm | 72.04 | -
| 13104 | Head Output Power nominal | dBm | 63 | -
| 13105 | Body Output Margin (fault peak power - norminal peak power) | dB | - | 1.5
| 13106 | Body Gain (output-input) | dB | 71.54 | 72.54
| 13107 | Head Delta Gain | dB | 62.44 | 63.64
| 13108 | Seq1 Body output power variation | % | - | 2
| 13109 | Seq2 Body output power variation | % | - | 2
| 13110 | Seq3 Body output power variation | % | - | 2
| 13111 | Seq4 Body output power variation | % | - | 2
| 13112 | Seq5 Body output power variation | % | - | 2
| 13113 | Seq6 Body output power variation | % | - | 2
| 13114 | Seq7Body output power variation | % | - | 2
| 13115 | Seq8 Body output power variation | % | - | 2
| 13201 | Harmonic Output | dBc | - | -30
| 13204 | Unblanked Output Noise Power broad spectrum | dBm/Hz | - | -70
| 13205 | Gain non-Linearity, -40 to 0 dBm input – Body Forward | dB | 0 | 0.8
| 13206 | Differential Gain, -40 to –3 dBm input– Body Forward | dB/dB | -0.1 | 0.1
| 13207 | Differential Gain, -3 to –1 dBm input - Body Forward | dB/dB | -0.2 | 0.1
| 13208 | Differential Gain, -1 to 0dBm input - Body Forward | dB/dB | -0.3 | 0.1
| 13209 | Phase non-Linearity, -40 to 0 dBm input – Body Forward | deg | 0 | 6
| 13210 | Differential Phase, -40 to – 3 dBm - Body Forward | deg/dB | -0.5 | 0.5
| 13211 | Differential Phase, -3 to – 1 dBm - Body Forward | deg/dB | -1.75 | 0.5
| 13212 | Differential Phase, -1 to – 0 dBm - Body Forward | deg/dB | -3 | 0.5
| 13213 | Gain non-Linearity, -40 to 0 dBm input – Body Reverse | dB | 0 | 0.8
| 13214 | Differential Gain, -40 to –3 dBm input– Body Reverse | dB/dB | -0.1 | 0.1
| 13215 | Differential Gain, -3 to –1 dBm input - Body Reverse | dB/dB | -0.2 | 0.1
| 13216 | Differential Gain, -1 to 0dBm input - Body Reverse | dB/dB | -0.3 | 0.1
| 13217 | Phase non-Linearity, -40 to 0 dBm input – Body Reverse | deg | 0 | 6
| 13218 | Differential Phase, -40 to – 3 dBm - Body Reverse | deg/dB | -0.5 | 0.5
| 13219 | Differential Phase, -3 to – 1 dBm - Body Reverse | deg/dB | -1.75 | 0.5
| 13220 | Differential Phase, -1 to – 0 dBm - Body Reverse | deg/dB | -3 | 0.5
| 13301 | Body Gain magnitude variation | dB | 0 | 0.35
| 13302 | Magnitude – Body 1 (Low Average Power) | dB | 0 | 0.09
| 13303 | Phase – Body 1(Low Average Power) | deg | 0 | 1.5
noise blanked | 13202 | Coherent noise (narrow spectrum) | dBm | - | -143 | ATE
| 13203 | Random noise (broad spectrum) | dBm/Hz | - | -160